Outline
- Motivation & Introduction
- Heating
- The Grand Master Plan
- Area Measurements
- Oscillating CO oxidation
- Sensitivity Measurements
- Time-Of-Flight Mass Spectrometry
- Summary

---
Introduction & Motivation
Many kinds of experimental catalysts needs to be tested in an early state of development, the μ-reactor provides an excellent platform for such tests.
Advantages includes:
* High sensitivity
* Low gas consumption
* High safety level
* Low thermal mass

---
μ-reactors
Photo: Thomas Pedersen

---
Experimental Setup
Photo: Mathias Kjærgaard Christensen

---
Experimental Setup
Photo: Mathias Kjærgaard Christensen

---
A Typical Measurement
CO oxidation on Pd thin film

---
General Measurement
Strategy

---
General Measurement
Strategy
The Grand Master Plan:
Thin film measurements
Nanoparticle measurements
Publish

---
Heating

---
Heating

---
Temperature profiles

---
Heating V2.0
CAD Drawing: Thomas Pedersen
Photo: Mathias Kjærgaard Christensen

---
Measuring surface area

---
Measuring surface area
We need a way of measuring the active area of the catalysts in the reactor

---
Area Measurements
Measurement principle:
CO-titration
Dose oxygen
Evacuate reactor
Dose CO
Watch for CO2
Limitations
Works only for Pt
Not very sensitive

---
Area Measurements
Calibration experiments performed on thin films
Easy cleaning
Known surface area
Stable system allows for lots of experimentation
Sensitivity....
Reproducebility....

---
Area Measurements
Measurement program:
Heating to 420◦ C in 1.0 bar O2
Heating to 300◦ C in 1.0 bar He
Heating to 300◦ C in 0.5 bar CO
Heating to 350◦ C in 1.0 bar O2
Heating to 300◦ C in 0.5 bar CO
Heating to 250◦ C in 1.0 bar O2
Heating to 300◦ C in 0.5 bar CO
Heating to 150◦ C in 1.0 bar O2
Heating to 300◦ C in 0.5 bar CO
According to the Grand Master Plan, we should now measure on nanoparticles.
The chosen sample has a nominal coverage of 10% of 5nm particles.
Structural modiﬁcation of Pt model systems under high pressure CO annealing
D.N. McCarthy, et. al C.E. Strebel, T.P. Johansson, A. den Dunnen, A. Nierhoff, J.H.Nielsen, and Ib Chorkendorff, 2012

---
Area Measurements
Nanoparticle sample
Thin-film sample

---
Area Measurements

---
Area Measurements
The Grand Master Plan:
Thin film measurements
Nanoparticle measurements
Publish
Establish new characterization tool for the reactors

---
Oscillating CO Oxidation
The story starts a dark and stormy saturday evening....
Experiments for a bachelors thesis a was done
and the setup was idling for the rest of the
weekend
An idling sample of size-selected nano-clusters is
the perfect time for random experiments!
Tried to run a classic CO oxidation light-off
experiment in extremely high oxygen
concentration
a
Mikkel Rykær Kraglund, Reactivity of mass
selected nanoparticles measured in microreactors

---
Oscillating Reactions
Data with unusually low quality called for an explanation

---
Oscillating Reactions

---
Oscillating Reactions
A more systematic study was initialized

---
Oscillating Reactions
Oscillations self sustained for several days, with steadily increasing period.

---
Oscillating Reactions

---
Oscillating Reactions
The Grand Master Plan:
Thin film measurements
Nanoparticle measurements
Publish Submit

---
Sensitivity Measurements
Observation.The sample loading for the oscillation experiments was very low.
Question. How small amounts of catalyst can be measured in the μ-reactor setup?

---
Sensitivity Measurements
Thomas Pedersen produced a wafer with Pt spots of different sizes. Six of these have so far been successfully measured.
#	Radius / μm	Area / μm2	Reactor coverage
N5	3.5	38.5	4.9 x 10-7
N8	15	707	9.0 x 10-6
N9	25	1960	2.5 x 10-5
N10	50	7850	1.0 x 10-4
N11	250	196000	2.5 x 10-3

---
Sensitivity Measurements
#	Radius / μm	Area / μm2	Reactor coverage
N5	3.5	38.5	4.9 x 10-7
N8	15	707	9.0 x 10-6
N9	25	1960	2.5 x 10-5
N10	50	7850	1.0 x 10-4
N11	250	196000	2.5 x 10-3

---
Sensitivity Measurements
Background correction
To do quantitative analysis, it is necessary to subtract the CO2 background.

---
Sensitivity Measurements
Background correction
To do quantitative analysis, it is necessary to subtract the CO2 background.
Background turns out to be dominated by oxygen:
ICO2 ~ 0.0082 x IO2

---
Sensitivity Measurements
Mass Spec Calibration
How do we get from MS-output to actual turn over frequency?
Use the approximately known total flow and calculate from the total SEM current.

---
Sensitivity Measurements

---
Sensitivity Measurements

---
Sensitivity Measurements

---
Sensitivity Measurements
- CO oxidation can be extremely efficient on Pt
- Even minute amounts of metal (espcially Pt) contamination must be avoided in the μ-reactors
The Grand Master Plan:
Thin film measurements
Nanoparticle measurements
Publish

---
Time-Of-Flight Mass Spectrometry

---
Time-Of-Flight Mass Spectrometry
Drawing:
Jordan TOF Products, Inc
Drawing: Thomas Andersen

---
Time-Of-Flight Mass Spectrometry

---
Time-Of-Flight Mass Spectrometry
Mass resolution much superior to a QMS allows for separation of OH and NH3

---
Time-Of-Flight Mass
Spectrometry
The Grand Master Plan:
Thin film measurements
Nanoparticle measurements
Thin film measurements
Publish

---
Conclusions